3i Spinning Disk Confocal

Yokogawa CSU-W1 T1 Spinning Disk with 50um pinhole disk

17mm x 16mm FOV

Prime 95B Back Illuminated sCMOS camera

1200x1200 pixels, 11x11um square pixels, 18.7mm FOV

95% Quantum Efficiency

Speed 0 = 12 bit/82 fps or 12msec/image - Gain only works in 12-bit mode

Speed 1 = 16-bit/41 fps or 24msec/image

Note: Binning on a CMOS camera does not increase imaging speed

CMOS – 4-pixel binning – 4 readouts

CCD – 4-pixel binning – 1 readout

To Crop – use Marquee tool to mark area and then select camera to update number of pixels

Lasers:	445nm – 100mW

488nm – 150mW

514nm – 150mW

561nm – 150mW

637nm – 140mW

Multi-Pass Excitation Dichroic/Single Emitter	Laser Position		Emission Filter Wheel

c445 		440/514/561			2 – 445nm		4 - 482/35 = 464.5-499.5nm BP

c488		405/488/561/640		3 – 488nm		5 - 525/30 = 510.0-540.0nm BP

c515		440/514/561			4 – 514nm		6 - 542/27 = 528.5-555.5nm BP

c561		405/488/561/640		5 – 561nm		7 - 617/73 = 580.5-653.5nm BP

c640		405/488/561/640		6 – 637nm		8 - 692/40 = 672.0-712.0nm BP

Quad-Pass Ex Dichroic/Quad BP Emitter	Laser Position		Emission Filter Wheel

q488		405/488/561/640		3 – 488nm		1 – 440/521/607/700

q561		405/488/561/640		5 – 561nm		1 – 440/521/607/700

q640		405/488/561/640		6 – 637nm		1 – 440/521/607/700

Triple-Pass Ex Dichroic/Tri BP Emitter		Laser Position		Emission Filter Wheel

t445		440/514/561			2 – 445nm		2 – 475/543/702

t515		440/514/561			4 – 514nm		2 – 475/543/702

t640		440/514/561			5 – 561nm		2 – 475/543/702