Stellaris 5 Start-Up Procedure

- Turn Power Button (#1) to the ON (“I”) position.

- Turn Laser Button (#2) to the ON (“I”) position.

DO NOT TURN THE KEY.  Leave the key in the vertical position.

- Turn the Computer (#3) ON.  The computer will take several minutes to initialize.

- Log into the software and double click on the LASX icon.

Select OK to initialize the Leica Application Suite X

Configuration: machine.xlhw

Microscope: DM18

- Select the Laser Overview icon to turn the laser on.

Change the Power State for each laser from OFF to ON.

Diode 405 (UV Light)

Diode 448 (Visible)

WLL (Supercon) 485-685

The WLL will take a few seconds to recognize and turn on.

Leave the WLL at 85% Maximum Power.

Stellaris 5 Shut-Down Procedure

- Exit the LASX software and shut down the computer.

- Turn the Laser Button (#2) to the OFF (“O”) position.

- Turn the Power Button (#1) to the OFF (“O”) position.